## Tabelle1
NOD-P1-NP Quantified using NIST 612 as a primary reference Material
[µg/g] | Mg | Ca | Sc | V | Co | Cu | Ga | As | Rb | Sr | Y | Mo | Cd | Sb | Cs | Ba | La | Ce | Pr | Nd | Sm | Eu | Gd | Tb | Dy | Ho | Er | Tm | Yb | Lu | W | Pb | Th | U
Result 1 | 18633.3333 | 19166.6667 | 8.13 | 411.6333 | 2076.6667 | 10933.3333 | 31.1433 | 117.2667 | 30.0367 | 519.5667 | 67.7433 | 706.9333 | 16.1367 | 61.87 | 2.15 | 2220 | 86.41 | 263.1667 | 25.1833 | 106.8333 | 26.4 | 6.53 | 23.7233 | 3.7233 | 20.1233 | 4.0933 | 11.1633 | 1.58 | 10.85 | 1.5333 | 67.8133 | 618.1 | 13.55 | 3.94
Result 2 | 18600 | 18900 | 7.9767 | 409.4 | 2073.3333 | 10900 | 30.92 | 114.3667 | 29.47 | 512.0667 | 66.01 | 699.3333 | 16.0633 | 61.5533 | 2.1567 | 2173.3333 | 84.39 | 260.3667 | 24.9533 | 103.8667 | 25.0933 | 6.2333 | 23.45 | 3.6767 | 19.62 | 3.8733 | 10.48 | 1.4733 | 10.3567 | 1.4867 | 65.7667 | 609.4 | 12.9933 | 3.7367
Result 3 | 18466.6667 | 18700 | 7.7667 | 409.4333 | 2093.3333 | 10933.3333 | 30.5267 | 118.0333 | 30.0167 | 500.1667 | 64.6467 | 713.3333 | 15.89 | 61.55 | 2.1867 | 2136.6667 | 83.2667 | 258.6667 | 24.1333 | 103 | 25.1833 | 6.2933 | 22.4667 | 3.5667 | 19.68 | 3.8367 | 10.2967 | 1.46 | 10.3933 | 1.4633 | 67.18 | 618.0667 | 12.94 | 3.7367
Result 4 | 18200 | 18866.6667 | 7.9667 | 399.9333 | 2106.6667 | 10766.6667 | 30.31 | 118 | 29.3833 | 500.3 | 64.0933 | 692.2 | 15.56 | 59.8967 | 2.1533 | 2086.6667 | 82.7167 | 252 | 23.69 | 103.5 | 25.06 | 5.9867 | 22.5633 | 3.6133 | 19.33 | 3.7733 | 10.2967 | 1.48 | 10.2467 | 1.45 | 65.47 | 606.4333 | 12.8833 | 3.7567
Result 5 | 18166.6667 | 19066.6667 | 7.8967 | 405.7 | 2090 | 10800 | 30.5667 | 115.4667 | 29.6233 | 503.5 | 64.4267 | 693.8 | 15.52 | 60.51 | 2.2267 | 2153.3333 | 83.8467 | 258.5667 | 24.3467 | 104.1333 | 24.87 | 6.3167 | 22.1 | 3.67 | 20.1833 | 3.8567 | 10.59 | 1.5467 | 10.6667 | 1.45 | 68.87 | 615.7333 | 13.0267 | 3.84
Result 6 | 18800 | 19566.6667 | 8.1133 | 415.4667 | 2080 | 10900 | 31.2133 | 116.8667 | 30.5 | 512.4333 | 67.0567 | 713.4333 | 16.0833 | 61.2067 | 2.23 | 2203.3333 | 87.8433 | 268.1333 | 25.4367 | 110.4 | 25.47 | 6.5467 | 23.66 | 3.7333 | 20.4533 | 4.0667 | 10.8233 | 1.59 | 11.0367 | 1.56 | 70.11 | 635.5 | 14.11 | 3.9767
Result 7 | 18733.3333 | 19300 | 8.14 | 410 | 2066.6667 | 10866.6667 | 31.04 | 115.5 | 29.3467 | 515.4333 | 66.5 | 688.9333 | 15.8467 | 59.92 | 2.1 | 2150 | 84.6033 | 262 | 24.4867 | 106.3333 | 25.2867 | 6.3333 | 22.9 | 3.5567 | 19.9533 | 3.8967 | 10.7833 | 1.5633 | 10.7267 | 1.5067 | 67.2133 | 610.8333 | 13.3667 | 3.8333
Average | 18514.2857 | 19080.9524 | 7.9986 | 408.7952 | 2083.8095 | 10871.4286 | 30.8171 | 116.5 | 29.7681 | 509.0667 | 65.7824 | 701.1381 | 15.8714 | 60.9295 | 2.1719 | 2160.4762 | 84.7252 | 260.4143 | 24.6043 | 105.4381 | 25.3376 | 6.32 | 22.9805 | 3.6486 | 19.9062 | 3.9138 | 10.6333 | 1.5276 | 10.611 | 1.4929 | 67.489 | 616.2952 | 13.2671 | 3.8314
Standard Error | 94.2408 | 110.7591 | 0.052578 | 1.8446 | 5.1654 | 24.5904 | 0.13165 | 0.53323 | 0.16221 | 2.9203 | 0.53489 | 3.8404 | 0.094306 | 0.30851 | 0.01747 | 16.7052 | 0.68351 | 1.8705 | 0.2334 | 0.99092 | 0.19074 | 0.071548 | 0.24171 | 0.026912 | 0.14562 | 0.045376 | 0.11868 | 0.020737 | 0.10906 | 0.016125 | 0.61907 | 3.6193 | 0.16814 | 0.03668
Expansion Factor [k] | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Expanded Uncertainty | 188.4816 | 221.5182 | 0.10516 | 3.6891 | 10.3309 | 49.1807 | 0.2633 | 1.0665 | 0.32441 | 5.8405 | 1.0698 | 7.6807 | 0.18861 | 0.61701 | 0.034941 | 33.4103 | 1.367 | 3.7409 | 0.46679 | 1.9818 | 0.38148 | 0.1431 | 0.48342 | 0.053824 | 0.29124 | 0.090751 | 0.23736 | 0.041473 | 0.21813 | 0.03225 | 1.2381 | 7.2386 | 0.33629 | 0.07336
Relative expanded Uncertainty [95% CL] | 1.018 | 1.1609 | 1.3147 | 0.90244 | 0.49577 | 0.45239 | 0.85439 | 0.91542 | 1.0898 | 1.1473 | 1.6262 | 1.0955 | 1.1884 | 1.0127 | 1.6088 | 1.5464 | 1.6135 | 1.4365 | 1.8972 | 1.8796 | 1.5056 | 2.2642 | 2.1036 | 1.4752 | 1.4631 | 2.3188 | 2.2322 | 2.7149 | 2.0557 | 2.1603 | 1.8346 | 1.1745 | 2.5348 | 1.9147
Horwitz RSD [%] | 3.6456 | 3.6291 | 11.6984 | 6.4714 | 5.0646 | 3.9497 | 9.5492 | 7.8171 | 9.5991 | 6.2612 | 8.5193 | 5.9667 | 10.5521 | 8.6181 | 14.2343 | 5.0371 | 8.2009 | 6.9258 | 9.8783 | 7.9354 | 9.8347 | 12.1205 | 9.9803 | 13.1653 | 10.1984 | 13.027 | 11.2077 | 15.0084 | 11.2112 | 15.0605 | 8.4865 | 6.0837 | 10.8406 | 13.0688
Horwitz Ratio [HORRAT] | 0.27925 | 0.3199 | 0.11238 | 0.13945 | 0.09789 | 0.11454 | 0.089472 | 0.1171 | 0.11353 | 0.18324 | 0.19089 | 0.1836 | 0.11262 | 0.1175 | 0.11302 | 0.30701 | 0.19674 | 0.20742 | 0.19206 | 0.23687 | 0.15309 | 0.1868 | 0.21077 | 0.11205 | 0.14346 | 0.178 | 0.19917 | 0.18089 | 0.18336 | 0.14344 | 0.21617 | 0.19306 | 0.23382 | 0.14651
The results listed above are the average of 3 measurements performed in 7 designated areas throughout the pellet's surface
The total number of measurements is 21
HORRAT of ≤ 1 is set as a benchmark for a usable reproducibility
These data are intended to illustrate an achievable reproducibiity. For calibration and validation purposes please use the values from the product information sheet
HORRAT = Measured Relative Standard Deviation / Calculated Relative Standard Deviation using the Horwitz equation
Horwitz, W., Albert, R. (1995). Precision in analytical measurements: Expected values and
and consequences in geochemical analyses. Fresenius Journal of Analytcal Chemistry.
351:507-513
Analytical setup
[11 empty] | ICP-MS |  | AttomES HR-ICP-MS, Nu Instruments
[11 empty] | Laser Ablation |  | Resonetics RESOlution M-50-LR 193nm eximer laser.
[13 empty] | Spot size | 48um
[13 empty] | Frequency | 5 Hz
[13 empty] | Fluence | 2.5 J/cm3
[13 empty] | Ablation time | 40 s
[13 empty] | Gas blank | 13 s
[13 empty] | Scan mode | Linked scan
[13 empty] | Total cycles | 550
[11 empty] | Quantification |  | NIST610
[11 empty] | Internal standard |  | Na
[11 empty] | Data evaluation |  | Iolite 4
Data were acquired at the University of Bergen, Norway.
We would like to thank Siv Hjorth Dundas for providing it.
Copyright © 2021 myStandards GmbH. All Rights Reserved.