## Tabelle1
NOD-A1-NP Quantified using NIST 610 as a primary reference Material
[µg/g] | Mg | P | Ca | Sc | Ti | V | Cr | Fe | Co | Ni | Cu | Zn | Ga | As | Rb | Sr | Mo | Cd | Sb | Ba | Pr | Nd | Sm | Eu | Gd | Dy | Er | Yb | W | Pb | Th | U
Result 1 | 22066.6667 | 4790 | 84633.3333 | 8.67 | 2003.3333 | 458.2333 | 17.65 | 91033.3333 | 2606.6667 | 5163.3333 | 1001.3333 | 550.9667 | 14.17 | 326.9 | 12.8333 | 1170 | 307.9333 | 4.7467 | 35.4533 | 1163.3333 | 17.8733 | 74.3633 | 16.7467 | 3.8933 | 17.6867 | 17.0067 | 10.2833 | 9.4833 | 67.18 | 1036.6667 | 17.5333 | 5.7367
Result 2 | 21833.3333 | 4880 | 86300 | 8.8967 | 2013.3333 | 452.0333 | 18.3533 | 90333.3333 | 2623.3333 | 5053.3333 | 983.6667 | 529.4667 | 13.9967 | 325.4667 | 12.6633 | 1186.6667 | 307.4333 | 4.64 | 34.79 | 1153.3333 | 17.9767 | 74.21 | 16.0867 | 3.8133 | 18.0767 | 16.99 | 10.4533 | 9.9 | 68.27 | 1053.3333 | 18.3033 | 5.7333
Result 3 | 22600 | 4863.3333 | 87000 | 9.03 | 2043.3333 | 468.1667 | 18.33 | 94600 | 2633.3333 | 5220 | 1017.6667 | 552.0333 | 14.1 | 331.1667 | 12.6667 | 1223.3333 | 316.7 | 4.64 | 35.4433 | 1176.6667 | 17.9667 | 76.8333 | 17.2733 | 3.9333 | 18.1233 | 17.39 | 10.6167 | 10.2033 | 69.8567 | 1033.3333 | 18.7 | 5.8
Result 4 | 23000 | 4990 | 90433.3333 | 9.1033 | 2116.6667 | 473.3667 | 18.98 | 94500 | 2750 | 5440 | 1047.3333 | 580.4 | 14.66 | 342.8333 | 13.6167 | 1236.6667 | 327.9667 | 4.9333 | 37.34 | 1216.6667 | 18.0767 | 78.2233 | 17.3233 | 4.01 | 18.1767 | 17.85 | 10.9367 | 10.1967 | 71.7233 | 1080 | 17.8933 | 5.9167
Result 5 | 22300 | 4883.3333 | 85600 | 8.7133 | 1993.3333 | 462.8 | 17.7033 | 89800 | 2656.6667 | 5183.3333 | 1012.6667 | 560.1333 | 13.58 | 327.6333 | 12.62 | 1170 | 312.5333 | 4.6333 | 36.3967 | 1150 | 17.7733 | 74.21 | 16.0567 | 3.8867 | 17.0433 | 17.0267 | 10.1667 | 9.43 | 69.9533 | 1033.3333 | 17.2533 | 5.6567
Result 6 | 21933.3333 | 4766.6667 | 85066.6667 | 8.6067 | 2000 | 454.6333 | 18.7833 | 89100 | 2586.6667 | 5156.6667 | 996.3333 | 544.8667 | 14.16 | 319.2 | 12.6267 | 1180 | 307.4333 | 4.54 | 35.31 | 1163.3333 | 17.78 | 74.3567 | 16.33 | 3.8667 | 17.7 | 16.8733 | 10.41 | 9.7967 | 69.0733 | 1036.6667 | 18.1467 | 5.7533
Result 7 | 22733.3333 | 4833.3333 | 86966.6667 | 9.06 | 2040 | 461.3667 | 18.38 | 91766.6667 | 2630 | 5200 | 1017 | 562 | 13.72 | 331.7333 | 12.8333 | 1190 | 312.3667 | 4.7167 | 35.5 | 1173.3333 | 17.6367 | 75.9767 | 16.87 | 3.9233 | 17.5467 | 17.4933 | 10.46 | 9.5767 | 69 | 1050 | 17.3333 | 5.8
Average | 22352.381 | 4858.0952 | 86571.4286 | 8.8686 | 2030 | 461.5143 | 18.3114 | 91590.4762 | 2640.9524 | 5202.381 | 1010.8571 | 554.2667 | 14.0552 | 329.2762 | 12.8371 | 1193.8095 | 313.1952 | 4.6929 | 35.7476 | 1170.9524 | 17.869 | 75.4533 | 16.6695 | 3.9038 | 17.7648 | 17.2329 | 10.4752 | 9.7981 | 69.2938 | 1046.1905 | 17.8805 | 5.771
Standard Error | 165.8483 | 27.6683 | 728.5351 | 0.077264 | 16.1998 | 2.8217 | 0.18796 | 828.8359 | 19.9697 | 44.4467 | 7.6514 | 5.9691 | 0.13216 | 2.7507 | 0.13438 | 9.8783 | 2.7847 | 0.047255 | 0.32019 | 8.4426 | 0.056763 | 0.60423 | 0.19933 | 0.023145 | 0.15227 | 0.1341 | 0.09385 | 0.12121 | 0.54129 | 6.4006 | 0.20323 | 0.030427
Expansion Factor [k] | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Expanded Uncertainty | 331.6967 | 55.3365 | 1457.0702 | 0.15453 | 32.3996 | 5.6434 | 0.37591 | 1657.6719 | 39.9395 | 88.8934 | 15.3028 | 11.9381 | 0.26431 | 5.5015 | 0.26876 | 19.7566 | 5.5694 | 0.09451 | 0.64039 | 16.8852 | 0.11353 | 1.2085 | 0.39866 | 0.046289 | 0.30454 | 0.26821 | 0.1877 | 0.24242 | 1.0826 | 12.8012 | 0.40645 | 0.060853
Relative expanded Uncertainty [95% CL] | 1.4839 | 1.1391 | 1.6831 | 1.7424 | 1.596 | 1.2228 | 2.0529 | 1.8099 | 1.5123 | 1.7087 | 1.5138 | 2.1539 | 1.8805 | 1.6708 | 2.0936 | 1.6549 | 1.7782 | 2.0139 | 1.7914 | 1.442 | 0.63533 | 1.6016 | 2.3916 | 1.1857 | 1.7143 | 1.5564 | 1.7918 | 2.4742 | 1.5623 | 1.2236 | 2.2732 | 1.0545
Horwitz RSD [%] | 3.5437 | 4.4588 | 2.8904 | 11.518 | 5.0845 | 6.3543 | 10.3274 | 2.866 | 4.8871 | 4.4131 | 5.6471 | 6.1816 | 10.7468 | 6.6856 | 10.8945 | 5.5075 | 6.7361 | 12.6759 | 9.3383 | 5.5235 | 10.3655 | 8.3453 | 10.4744 | 13.032 | 10.3746 | 10.4222 | 11.233 | 11.3465 | 8.4529 | 5.618 | 10.3645 | 12.2875
Horwitz Ratio [HORRAT] | 0.41876 | 0.25546 | 0.58231 | 0.15128 | 0.3139 | 0.19244 | 0.19878 | 0.6315 | 0.30945 | 0.38719 | 0.26808 | 0.34843 | 0.17499 | 0.24991 | 0.19217 | 0.30049 | 0.26399 | 0.15888 | 0.19184 | 0.26107 | 0.061293 | 0.19192 | 0.22832 | 0.090987 | 0.16524 | 0.14933 | 0.15952 | 0.21805 | 0.18482 | 0.2178 | 0.21932 | 0.085817
The results listed above are the average of 3 measurements performed in 7 designated areas throughout the pellet's surface
The total number of measurements is 21
HORRAT of ≤ 1 is set as a benchmark for a usable reproducibility
These data are intended to illustrate an achievable reproducibiity. For calibration and validation purposes please use the values from the product information sheet
HORRAT = Measured Relative Standard Deviation / Calculated Relative Standard Deviation using the Horwitz equation
Horwitz, W., Albert, R. (1995). Precision in analytical measurements: Expected values and
and consequences in geochemical analyses. Fresenius Journal of Analytcal Chemistry.
351:507-513
Analytical setup
[11 empty] | ICP-MS |  | AttomES HR-ICP-MS, Nu Instruments
[11 empty] | Laser Ablation |  | Resonetics RESOlution M-50-LR 193nm eximer laser.
[13 empty] | Spot size | 48um
[13 empty] | Frequency | 5 Hz
[13 empty] | Fluence | 2.5 J/cm3
[13 empty] | Ablation time | 40 s
[13 empty] | Gas blank | 13 s
[13 empty] | Scan mode | Linked scan
[13 empty] | Total cycles | 550
[11 empty] | Quantification |  | NIST610
[11 empty] | Internal standard |  | Na
[11 empty] | Data evaluation |  | Iolite 4
Data were acquired at the University of Bergen, Norway.
We would like to thank Siv Hjorth Dundas for providing it.
Copyright © 2021 myStandards GmbH. All Rights Reserved.