## Tabelle1
NOD-P1-NP Quantified using NIST 610 as a primary reference Material
[µg/g] | Mg | P | Ca | Sc | V | Fe | Co | Ni | Cu | Ga | As | Rb | Sr | Mo | Cd | Sb | Cs | Ba | Sm | Eu | Gd | Dy | Er | Yb | W | Pb | Th | U
Result 1 | 16733.3333 | 2266.6667 | 19333.3333 | 8.16 | 415.2 | 58166.6667 | 2066.6667 | 13133.3333 | 12133.3333 | 31.5033 | 111.6333 | 30.5067 | 520.5333 | 686.7 | 16.7267 | 60.66 | 2.06 | 2166.6667 | 26.5967 | 6.2167 | 23.6367 | 19.7733 | 10.6433 | 9.69 | 59.9533 | 583.5333 | 11.07 | 3.7167
Result 2 | 16700 | 2300 | 18933.3333 | 8.0267 | 413.4 | 58333.3333 | 2066.6667 | 12566.6667 | 12166.6667 | 31.49 | 110 | 29.9933 | 515.5667 | 683.9 | 16.7133 | 60.8533 | 2.0867 | 2133.3333 | 25.3267 | 5.9967 | 23.46 | 19.6933 | 10.2367 | 9.4567 | 60.77 | 592.3667 | 11.5367 | 3.7533
Result 3 | 16566.6667 | 2366.6667 | 18733.3333 | 7.8033 | 413.9333 | 56066.6667 | 2100 | 12500 | 12166.6667 | 30.9767 | 112.6333 | 30.6033 | 501.9333 | 695.8667 | 16.5 | 60.72 | 2.1 | 2100 | 25.4433 | 5.9633 | 22.3033 | 19.59 | 9.96 | 9.4533 | 60.6067 | 594.6333 | 11.03 | 3.6333
Result 4 | 16333.3333 | 2333.3333 | 18966.6667 | 7.9333 | 405.5667 | 56100 | 2100 | 12533.3333 | 11966.6667 | 30.5667 | 111.1667 | 29.9833 | 502.7 | 678.1667 | 16.2567 | 58.9433 | 2.0633 | 2066.6667 | 25.3133 | 5.6733 | 22.2967 | 19.2467 | 9.7733 | 9.3067 | 58.4033 | 572.9333 | 11.62 | 3.5667
Result 5 | 16366.6667 | 2266.6667 | 19166.6667 | 7.8667 | 410.9333 | 57533.3333 | 2100 | 12500 | 12066.6667 | 30.98 | 110.2 | 30.12 | 506.2333 | 682.0667 | 16.2 | 59.5033 | 2.1467 | 2133.3333 | 25.14 | 5.9833 | 22.2333 | 20.1967 | 10.1167 | 9.6267 | 61.1567 | 579.0333 | 11.6767 | 3.65
Result 6 | 16866.6667 | 2366.6667 | 19700 | 8.1067 | 420.3667 | 56766.6667 | 2100 | 13033.3333 | 12166.6667 | 31.7833 | 111.9667 | 31.0033 | 515.6667 | 699.6667 | 16.7367 | 60.3267 | 2.1367 | 2133.3333 | 25.72 | 6.1733 | 23.4367 | 20.48 | 10.3533 | 9.7933 | 62.47 | 604.6 | 12.09 | 3.8133
Result 7 | 16800 | 2300 | 19500 | 8.14 | 415.4667 | 56733.3333 | 2033.3333 | 12866.6667 | 12100 | 31.3233 | 109.1 | 30.07 | 517.5 | 675.5333 | 16.5767 | 58.9467 | 2.0233 | 2133.3333 | 25.52 | 6.01 | 22.8333 | 19.7667 | 10.3133 | 9.69 | 60.7033 | 574.9333 | 12 | 3.64
Average | 16623.8095 | 2314.2857 | 19190.4762 | 8.0052 | 413.5524 | 57100 | 2080.9524 | 12733.3333 | 12109.5238 | 31.2319 | 110.9571 | 30.3257 | 511.4476 | 685.9857 | 16.53 | 59.9933 | 2.0881 | 2123.8095 | 25.58 | 6.0024 | 22.8857 | 19.821 | 10.1995 | 9.5738 | 60.5805 | 586.0048 | 11.5748 | 3.6819
Standard Error | 78.9194 | 16.031 | 129.2457 | 0.053004 | 1.7156 | 350.5098 | 9.9127 | 102.8689 | 27.9022 | 0.15613 | 0.46976 | 0.14674 | 2.8792 | 3.3644 | 0.084753 | 0.31853 | 0.016569 | 11.9839 | 0.18286 | 0.066439 | 0.23461 | 0.15276 | 0.107 | 0.065025 | 0.46515 | 4.3884 | 0.15521 | 0.031622
Expansion Factor [k] | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Expanded Uncertainty | 157.8388 | 32.0619 | 258.4915 | 0.10601 | 3.4311 | 701.0197 | 19.8254 | 205.7378 | 55.8044 | 0.31227 | 0.93952 | 0.29348 | 5.7585 | 6.7288 | 0.16951 | 0.63706 | 0.033138 | 23.9677 | 0.36572 | 0.13288 | 0.46922 | 0.30553 | 0.21399 | 0.13005 | 0.9303 | 8.7769 | 0.31041 | 0.063243
Relative expanded Uncertainty [95% CL] | 0.94947 | 1.3854 | 1.347 | 1.3242 | 0.82967 | 1.2277 | 0.95271 | 1.6157 | 0.46083 | 0.99983 | 0.84674 | 0.96777 | 1.1259 | 0.9809 | 1.0254 | 1.0619 | 1.587 | 1.1285 | 1.4297 | 2.2138 | 2.0503 | 1.5414 | 2.0981 | 1.3584 | 1.5356 | 1.4978 | 2.6818 | 1.7177
Horwitz RSD [%] | 3.7052 | 4.9852 | 3.626 | 11.6969 | 6.4601 | 3.0772 | 5.0656 | 3.8569 | 3.8861 | 9.53 | 7.8747 | 9.5723 | 6.2568 | 5.9864 | 10.4877 | 8.6382 | 14.3188 | 5.0501 | 9.8207 | 12.215 | 9.9866 | 10.205 | 11.2782 | 11.3861 | 8.6256 | 6.13 | 11.0655 | 13.1473
Horwitz Ratio [HORRAT] | 0.25626 | 0.2779 | 0.37148 | 0.11321 | 0.12843 | 0.39897 | 0.18807 | 0.41892 | 0.11858 | 0.10491 | 0.10753 | 0.1011 | 0.17995 | 0.16386 | 0.097776 | 0.12293 | 0.11083 | 0.22347 | 0.14558 | 0.18123 | 0.2053 | 0.15105 | 0.18603 | 0.1193 | 0.17803 | 0.24433 | 0.24236 | 0.13065
The results listed above are the average of 3 measurements performed in 7 designated areas throughout the pellet's surface
The total number of measurements is 21
HORRAT of ≤ 1 is set as a benchmark for a usable reproducibility
These data are intended to illustrate an achievable reproducibiity. For calibration and validation purposes please use the values from the product information sheet
HORRAT = Measured Relative Standard Deviation / Calculated Relative Standard Deviation using the Horwitz equation
Horwitz, W., Albert, R. (1995). Precision in analytical measurements: Expected values and
and consequences in geochemical analyses. Fresenius Journal of Analytcal Chemistry.
351:507-513
Analytical setup
[11 empty] | ICP-MS |  | AttomES HR-ICP-MS, Nu Instruments
[11 empty] | Laser Ablation |  | Resonetics RESOlution M-50-LR 193nm eximer laser.
[13 empty] | Spot size | 48um
[13 empty] | Frequency | 5 Hz
[13 empty] | Fluence | 2.5 J/cm3
[13 empty] | Ablation time | 40 s
[13 empty] | Gas blank | 13 s
[13 empty] | Scan mode | Linked scan
[13 empty] | Total cycles | 550
[11 empty] | Quantification |  | NIST610
[11 empty] | Internal standard |  | Na
[11 empty] | Data evaluation |  | Iolite 4
Data were acquired at the University of Bergen, Norway.
We would like to thank Siv Hjorth Dundas for providing it.
Copyright © 2021 myStandards GmbH. All Rights Reserved.