## Characterisation
Characterisation
Analyte | Strontium - Sr | µg/g
| Method 1 | Method 2 | Method 3 | Method 4 | Certified Value | 2705.1415 | µg/g
Data Set Mean 1 | 2600 | 2795 | 2705.913 | 2621.3532 | Uncertainty component Characterisation | 12.5474 | µg/g
Data Set Mean 2 | 2690 | 2685 | 2663.6331 | 2663.6331
Data Set Mean 3 | 2600 | 2685 | 2621.3532 | 2663.6331 | Calculation of certified value and uncertainty component following Annex A.2.4 equation A.1 &
Data Set Mean 4 | 2780 | 2750 | 2705.913 | 2663.6331 | Annex A.2.5.3 equation A.4 in ISO Guide 35:2017
Data Set Mean 5 | 2615 | 2840 | 2705.913 | 2663.6331
Data Set Mean 6 | 2745 | 2830 | 2663.6331 | 2663.6331
Data Set Mean 7 | 2680 | 2840 | 2663.6331 | 2705.913
Data Set Mean 8 | 2810 | 2790 | 2663.6331 | 2790.4728
These data sets were acquired in accredited ISO 17025:2017 laboratories and as such are SI-traceable.
Approximate normal distribution of data set means was checked using XLSTAT, data are not shown here since the add-in needs to be installed.
EN ISO/IEC 17025:2017 (D/E): General requirements for the competence of testing and calibration laboratories
ISO Guide 35:2017 (E): Reference materials – Guidance for characterization and assessment of homogeneity and stability
Copyright © 2022 myStandards GmbH. All Rights Reserved.

## Homogeneity Test Single Pellet
Analyte |  | Sr | Apatite-NP 2020-07 B01-80
[µg/g]
Apatite-NP-2020-07-B01-80 1a | 1a | 2102.9353 |  | Zone 1 | Zone 2 | Zone 3 | Zone 4 | Zone 5 | Zone 6 | Zone 7 | Average | SD | RSD [%]
Apatite-NP-2020-07-B01-80 6a | 6a | 2107.8839 | Result 1 (a) | 2102.9353 | 2119.8305 | 2096.5516 | 2145.5504 | 2106.8876 | 2107.8839 | 2100.1256 | 2111.395 | 16.7794 | 0.79471
Apatite-NP-2020-07-B01-80 7a | 7a | 2100.1256 | Result 2 (b) | 2087.9862 | 2092.1626 | 2083.5473 | 2104.6937 | 2110.7049 | 2082.9166 | 2113.736 | 2096.5353 | 12.9716 | 0.61871
Apatite-NP-2020-07-B01-80 2a | 2a | 2119.8305 | Result 3 (c) | 2139.5145 | 2126.6516 | 2067.6847 | 2145.1968 | 2069.4183 | 2118.4659 | 2116.9735 | 2111.9865 | 31.4042 | 1.487
Apatite-NP-2020-07-B01-80 5a | 5a | 2106.8876
Apatite-NP-2020-07-B01-80 3a | 3a | 2096.5516 | Average | 2110.1454 | 2112.8816 | 2082.5945 | 2131.8136 | 2095.6703 | 2103.0888 | 2110.2784 | Summary |  |  |  |  | Student's T
Apatite-NP-2020-07-B01-80 4a | 4a | 2145.5504 | SD | 26.51 | 18.2644 | 14.457 | 23.4872 | 22.8149 | 18.2533 | 8.9403 | [5 empty] | 95% CI two-tailed
NIST610- |  | 518.1156 | RSD [%] | 1.2563 | 0.86443 | 0.69418 | 1.1017 | 1.0887 | 0.86793 | 0.42366 | [5 empty] | n | k
NIST610- |  | 514.1594 | [13 empty] | 1 | 12.71
NIST610- |  | 513.2624 | [9 empty] | Average | 2106.6389 | [µg/g] |  | 2 | 4.303
Apatite-NP-2020-07-B01-80 3b | 3b | 2083.5473 | Anova: One-way analysis of variance |  |  | α= | 0.05 |  |  |  |  | Standard error | 5.7927 | [µg/g] |  | 3 | 3.182
Apatite-NP-2020-07-B01-80 7b | 7b | 2113.736 | [9 empty] | Exp. Unc. | 11.5854 | [95% CI] |  | 4 | 2.776
Apatite-NP-2020-07-B01-80 2b | 2b | 2092.1626 | SUMMARY | [8 empty] | k | 2 | [-] |  | 5 | 2.571
Apatite-NP-2020-07-B01-80 5b | 5b | 2110.7049 | Groups | Count | Sum | Average | Variance |  |  |  |  | degrees of freedom | 20 | [-] |  | 6 | 2.447
Apatite-NP-2020-07-B01-80 6b | 6b | 2082.9166 | Zone 1 | 3 | 6330.4361 | 2110.1454 | 702.781 | [8 empty] | 7 | 2.365
Apatite-NP-2020-07-B01-80 1b | 1b | 2087.9862 | Zone 2 | 3 | 6338.6448 | 2112.8816 | 333.5882 |  |  |  |  | rel. Unc. RSD-% | 0.27497 | [%] |  | 8 | 2.306
Apatite-NP-2020-07-B01-80 4b | 4b | 2104.6937 | Zone 3 | 3 | 6247.7836 | 2082.5945 | 209.0052 |  |  |  |  | RSD Horwitz | 5.0563 | [%] |  | 9 | 2.262
NIST610- |  | 519.0895 | Zone 4 | 3 | 6395.4409 | 2131.8136 | 551.6503 |  |  |  |  | HORRAT | 0.054383 | [-] |  | 10 | 2.228
NIST610- |  | 512.0143 | Zone 5 | 3 | 6287.0108 | 2095.6703 | 520.5181 | [8 empty] | >10 | 2
NIST610- |  | 515.4558 | Zone 6 | 3 | 6309.2664 | 2103.0888 | 333.1844 |  |  |  | A HORRAT < 1 is great, between 1 and 2 acceptable and >2 unacceptable |  |  |  | Horwitz, W., Albert, R. (1995). Precision in analytical measurements: Expected values and
Apatite-NP-2020-07-B01-80 5c | 5c | 2069.4183 | Zone 7 | 3 | 6330.8351 | 2110.2784 | 79.9297 | [7 empty] | and consequences in geochemical analyses. Fresenius Journal of Analytical Chemistry.
Apatite-NP-2020-07-B01-80 7c | 7c | 2116.9735 | [12 empty] | 351:507-513
Apatite-NP-2020-07-B01-80 2c | 2c | 2126.6516 | [8 empty] | ASTM-E826-14 Homogeneity Test
Apatite-NP-2020-07-B01-80 3c | 3c | 2067.6847 | ANOVA
Apatite-NP-2020-07-B01-80 6c | 6c | 2118.4659 | Source of Variation | Sum of Squares (SS) | Degrees of freedom | Mean sum of squares | F | P-Value | F crit
Apatite-NP-2020-07-B01-80 1c | 1c | 2139.5145 | Between Groups | 4227.9726 | 6 | 704.6621 | 1.8064 | 0.16963 | 2.8477
Apatite-NP-2020-07-B01-80 4c | 4c | 2145.1968 | Within Groups | 5461.3138 | 14 | 390.0938 | [5 empty] | Number of Measurements per Zone (b) | 3
NIST610- |  | 511.1586 | [9 empty] | Number of Zones (t) | 7
NIST610- |  | 519.7377 | Total | 9689.2863 | 20 | [6 empty] | Degrees of Freedom (n) | 12
NIST610- |  | 513.6162
[12 empty] | SSt | 4227.9726
The analyses were always bracketed by 4 blocks of 3 |  |  | Anova: One-way analysis of variance | [8 empty] | SSb | 1073.0946
measurements of SRM NIST 610 for drift correction. | [11 empty] | SST | 9689.2863
The uppermost block is not shown here. |  |  | SUMMARY
|  |  | Groups | Count | Sum | Average | Variance |  |  |  |  | s | 19.1229
|  |  | Result 1 (a) | 7 | 14779.7649 | 2111.395 | 281.5473 |  |  |  |  | q | 4.95
|  |  | Result 2 (b) | 7 | 14675.7473 | 2096.5353 | 168.2615
|  |  | Result 3 (c) | 7 | 14783.9054 | 2111.9865 | 986.2232 |  |  |  |  | w | 54.651
[12 empty] | maximum(t')-minimum(t') | 49.2191
|  |  | ANOVA | [7 empty] | The homogeneity test passes if maximum(t') - minimum (t') is < than the "w" value
|  |  | Source of Variation | Sum of Squares (SS) | Degrees of freedom | Mean sum of squares | F | P-Value | F crit
|  |  | Between Groups | 1073.0946 | 2 | 536.5473 | 1.1209 | 0.34771 | 3.5546 |  | ASTM E826-14: Standard Practice for Testing Homogeneity of a Metal Lot or Batch in Solid
|  |  | Within Groups | 8616.1918 | 18 | 478.6773 |  |  |  |  | form by Spark Atomic Emission Spectroscopy. ASTM International, West Conshohocken, PA, 2014
www.astm.org
|  |  | Total | 9689.2863 | 20
This type of homogeneity Test was performed on a total of 10 Apatite-NP Nano-Pellets. LA-ICP-MS data quantified using NIST 610 and Ca as internal standard.
The amount of pellets needed varies with the amount of pellets produced.
Ablation pattern on the pellet surface
Copyright © 2022 myStandards GmbH. All Rights Reserved.

## Homogeneity- & Stability Test
| Analyte | Sr
[µg/g]
| Homogeneity | Result 1 | Result 2 | Result 3 | Result 4 | Result 5 | Result 6 | Result 7
| Apatite-NP-2020-07-B01-80 | 2110.1454 | 2112.8816 | 2082.5945 | 2131.8136 | 2095.6703 | 2103.0888 | 2110.2784
| Apatite-NP-2020-07-B01-22 | 2117.5963 | 2118.2785 | 2113.3519 | 2090.8318 | 2121.3176 | 2144.6614 | 2113.9106
| Apatite-NP-2020-07-B01-97 | 2119.4831 | 2092.2767 | 2136.2222 | 2125.4587 | 2143.2788 | 2122.0028 | 2105.24
| Apatite-NP-2020-07-B01-37 | 2110.0948 | 2099.5522 | 2112.8243 | 2114.5156 | 2133.9581 | 2121.6125 | 2108.0255
| Apatite-NP-2020-07-B01-117 | 2125.2175 | 2129.4602 | 2118.7775 | 2134.3097 | 2117.45 | 2111.2212 | 2120.3536
| Apatite-NP-2020-07-B01-68 | 2122.2104 | 2102.6983 | 2130.5629 | 2117.4995 | 2102.7853 | 2132.7907 | 2104.3664
| Apatite-NP-2020-07-B01-49 | 2145.2731 | 2148.1711 | 2149.7939 | 2149.2809 | 2144.0665 | 2141.6084 | 2137.4634
| Apatite-NP-2020-07-B01-11 | 2135.2951 | 2130.8946 | 2152.73 | 2127.5175 | 2142.6848 | 2123.0844 | 2134.1897
| Apatite-NP-2020-07-B01-89 | 2129.9694 | 2157.4317 | 2138.1069 | 2123.8535 | 2138.8222 | 2147.1517 | 2139.1092
| Apatite-NP-2020-07-B01-10 | 2110.1454 | 2127.4856 | 2129.0098 | 2148.3628 | 2137.4263 | 2158.7477 | 2140.8107
| Stability | Result 1 | Result 2 | Result 3 | Result 4 | Result 5 | Result 6 | Result 7
| Apatite-NP-2020-07-B01-40 | 2161.0601 | 2156.4381 | 2127.0193 | 2138.4635 | 2170.353 | 2158.622 | 2145.4083
| Apatite-NP-2020-07-B01-96 | 2164.1633 | 2145.5632 | 2117.295 | 2143.7672 | 2151.5478 | 2122.228 | 2103.7546
| Anova: One-way analysis of variance |  |  | α= | 0.05
SUMMARY
| Groups | Count | Sum | Average | Variance
| Apatite-NP-2020-07-B01-80 | 7 | 14746.4725 | 2106.6389 | 234.8874
| Apatite-NP-2020-07-B01-22 | 7 | 14819.9481 | 2117.1354 | 248.8801
| Apatite-NP-2020-07-B01-97 | 7 | 14843.9622 | 2120.566 | 303.8886
| Apatite-NP-2020-07-B01-37 | 7 | 14800.5831 | 2114.369 | 119.4421
| Apatite-NP-2020-07-B01-117 | 7 | 14856.7897 | 2122.3985 | 61.0676
| Apatite-NP-2020-07-B01-68 | 7 | 14812.9134 | 2116.1305 | 170.2675
| Apatite-NP-2020-07-B01-49 | 7 | 15015.6573 | 2145.0939 | 20.0919
| Apatite-NP-2020-07-B01-11 | 7 | 14946.3961 | 2135.1994 | 98.1165
| Apatite-NP-2020-07-B01-89 | 7 | 14974.4446 | 2139.2064 | 119.6145
| Apatite-NP-2020-07-B01-10 | 7 | 14951.9883 | 2135.9983 | 247.549
ANOVA
| Source of Variation | Sum of Squares (SS) | Degrees of freedom | Mean Sum of Squares | F | P-Wert | F crit
| Between Groups | 10128.4576 | 9 | 1125.3842 | 6.9305 | 8.9333e-07 | 2.0401
| Within Groups | 9742.8307 | 60 | 162.3805
| Total | 19871.2882 | 69
| Pellet Evaluation | following ISO Guide 35:2017 Annex C1. p. 89f
|  | s2between | 137.572 | [µg^2/g^2]
|  | sbb | 11.7291 | [µg/g]
|  | sr | 12.7429 | [µg/g]
|  | Average | 2125.2736 | [µg/g]
|  | Combined Unc. Homogeneity | 17.3191 | [µg/g]
|  | rel. combined Unc. Homogeneity | 0.81491 | [%]
ISO Guide 35:2017 (E): Reference materials – Guidance for characterization and assessment of homogeneity and stability
| Pellet Evaluation | following ISO 13528:2015 Annex E2. p. 68f
|  | Usability (expected RSD Horwitz) | 5.0496 | [%]
|  | Usability decimal | 0.050496 | [-]
| Homogeneity | Average Homogeneity | 2125.2736 | [µg/g]
|  | Standard deviation of Averages | 12.6795 | [µg/g]
|  | sw | 12.7429 | [µg/g]
|  | ss | 11.7291 | [µg/g]
|  | σpt | 107.3171
|  | Check value homogeneity | 32.1951
| Stability | Average Stability | 2143.2631 | [µg/g]
|  | Absolute difference between Homogeneity Average | 17.9895 | [µg/g]
|  | Check value stability | 32.1951 |  | ISO 13528:2015 (E): Statistical methods for use in proficiency testing by interlaboratory comparison
Copyright © 2022 myStandards GmbH. All Rights Reserved.

## Value & Uncertainty Calculation
Certified Value and Uncertainty Calculation
Analyte | Sr
Certified Value | 2705.1415 | µg/g | Pellet Evaluation | following ISO Guide 35:2017 Annex C1. p. 89f
Uncertainty component Characterisation | 12.5474 | µg/g
Uncertainty component Homogeneity | 22.0446 | µg/g |  | s2between | 137.572 | [µg^2/g^2]
Uncertainty component Stability | 22.7056 | µg/g |  | sbb | 11.7291 | [µg/g]
|  |  |  | sr | 12.7429 | [µg/g]
Combined Uncertainty | 34.0433 | µg/g
expansion factor | 2 | [-] |  | Average | 2125.2736 | [µg/g]
|  |  |  | Combined Unc. Homogeneity | 17.3191 | [µg/g]
Combined expanded Uncertainty @ 95 % CL | 68.0865 | µg/g |  | rel. combined Unc. Homogeneity | 0.81491 | [%]
Relative deviation from certified value | 2.5169 | %
Horwitz relative deviation | 4.8695 | %
HORRAT | 0.51688 | [-] | Pellet Evaluation | following ISO 13528:2015 Annex E2. p. 68f
The quantified values on the Nano-Pellets are traceable to SRM NIST 610,
but due to matrix-effects during LA-ICP-MS a factor for the uncertainty contribution |  |  |  | Usability (exp. RSD Horwitz) | 5.0496 | [%]
from homogeneity and stability is calculated instead of using the quantified value. |  |  |  | Usability decimal | 0.050496 | [-]
The certified value is derived from SI-traceable data from accredited |  |  | Homogeneity | Average Homogeneity | 2125.2736 | [µg/g]
(ISO 17025:2017) laboratories. Therefore the measured uncertainty contribution is |  |  |  | Standard deviation of Averages | 12.6795 | [µg/g]
used instead of a factor. |  |  |  | sw | 12.7429 | [µg/g]
|  |  |  | ss | 11.7291 | [µg/g]
A HORRAT of < 1 up to 2 signifies a usable combined expanded Uncertainty. |  |  |  | σpt | 107.3171
HORRAT of >2 are unacceptable and the value becomes informative |  |  |  | Check value homogeneity | 32.1951
Ultimately, the decision whether an uncertainty is usable or not lies with the
manufacturer and the material's intended use. Meaning even a HORRAT > 2 |  |  | Stability | Average Stability | 2143.2631 | [µg/g]
could still be deemed usable. |  |  |  | Absolute difference between Homogeneity Average | 17.9895 | [µg/g]
|  |  |  | Check value stability | 32.1951
Horwitz, W., Albert, R. (1995). Precision in analytical measurements: Expected values and
and consequences in geochemical analyses. Fresenius Journal of Analytical Chemistry.
351:507-513
Copyright © 2022 myStandards GmbH. All Rights Reserved.